Phycocyanin Extraction Methods

25Apr24

- Filter water samples onto GF/F filters under low light. Record volume filtered.

- Freeze the GF/F filters in foil packets. Freezing is an important step- don’t skip.

- Thaw filters

- Under reduced light, cut the filters into small pieces and grind them in 5 mL of 50 mM phosphate buffer with a Teflon grinder for two minutes. Be sure to rinse the Teflon grinder with DI water and dry in between each sample

- Use 15 mL of phosphate buffer to rinse sample into a 50 ml centrifuge tube

- Cap the cientrifuge tubes and mix samples.

- Cover the samples with foil and leave them in the fridge for 2 hours

- After the 2 hours in the fridge, transfer the tubes to a dark cabinet for an additional 2 hours

- Filter 4 ml of the samples through a 25 mm size 1.6 µm pore GF/F filter. Save the filtrate

- Filter the filtrate through a 25 mm size 0.45 µm Millipore HNWP white nylon filter. Save the filtrate

- Filter the filtrate one last time through a 25 mm 0.2 µm Millipore GTTP filter. Save the filtrate.

- Transfer 2 ml of the filtrate to the Nova Biotech cuvette.

- Place the cuvette into the Turner Trilogy fluorometer using the PC orange module and run the sample. Do not use a standard curve.

- Record the RFU values for each sample

- Rinse and dry cuvette in between samples. Be super careful to not scratch the cuvette.

- Convert the RFU values into concentration using standard curve values. Adjust concentration for volume filtered and volume of phosphate buffer